# General License

# Questions and Answers

Regarding Radioactive Materials Distributed Under a General License

### NOTIFICATIONS TO THE DEPARTMENT OF NEW GL DEVICE

- What is the Vendor Report?

- When are the Vendor Reports due?

- What information is the distributor required to give to the general license?

### REGISTERED GENERAL LICENSEE OBLIGATIONS

- What devices require registration?

- What is the registrant obligated to do?

- What is involved in registering?

- How often do these devices need to be registered?

- How do registrants know they need to register these devices, and when do registrants need to register by?

- What information needs to be supplied to register a specified generally licensed device?

- When is the registrant obligated to contact the department?

- What other regulations is the registrant subjected to?

### REGISTRATION PROCESS

- How does a registrant receive a registration form?

- What will the registrant receive at the end of the registration process?

- Will the State of Washington Radioactive Material General License Registration have an expiration date?

### FEES

Current Annual Fees

### DISPOSAL AND TRANSFERS

- How may the registrant dispose of the registered device?

- May the registrant transfer the registered device to another general licensee?

### NOTIFICATIONS TO THE DEPARTMENT OF NEW GL DEVICE

- What is the Vendor Report? A report that tells the department of all the generally licensed devices transferred to a general license recipient or intermediate person, or devices received from a general licensee.

WAC 246-235-093  Manufacture, assembly or distribution of devices under general license.

WAC 246-235-093(4)(c) Report to the department all transfers of such devices to persons for use under the general license in WAC 246-233-020 and all receipts of devices from persons licensed under WAC 246-233-020.

(i) Such report shall  include:

(A) The identity of each general licensee by name and mailing address for the location of use; if there is no mailing address for the location of use, an alternative address for the general licensee shall be submitted along with information on the actual location of use;

(B) The name, title, and phone number of the person identified by the general licensee as having knowledge of and authority to take required actions to ensure compliance with the appropriate regulations and requirements;

(C) The date of transfer;

(D) The type, model number and serial number of device transferred; and

(E) The quantity and type of radioactive material contained in the device.

(ii) If one or more intermediate persons will temporarily possess the device at the intended place of use prior to its possession by the user, the report shall include identification of each intermediate person by name, address, contact, and relationship to the intended user.

(iii) For devices received from persons generally licensed under WAC 246-233-020, the report must include:

(A) The identity of the general licensee by name and address;

(B) The type, model number, and serial number of the device received;

(C) The date of receipt; and

(D) In the case of devices not initially transferred by the reporting licensee, the name of the manufacturer or initial transferor.

(iv) If the licensee makes changes to a device possessed by a person generally licensed under WAC 246-233-020, such that the label must be changed to update required information, the report must identify the general licensee, the device, and the changes to information on the device label.

(v) If no transfers have been made to or from persons generally licensed under WAC 246-233-020 during the reporting period, the report shall so indicate.

(vi) The report shall cover each calendar quarter, shall clearly indicate the period covered by the report, and shall be filed within thirty days of the end of the calendar quarter.

(vii) The report shall clearly identify the specific licensee submitting the report and include the license number of the specific licensee.

- When are the Vendor Reports due?

WAC 246-235-093(4)(c)(vi) The report shall cover each calendar quarter, shall clearly indicate the period covered by the report, and shall be filed within thirty days of the end of the calendar quarter.

- What information is the distributor required to give to the general license?

WAC 246-235-093(4) Each person licensed under subsection (1) of this section to distribute or initially transfer devices to generally licensed persons shall, prior to the transfer to the intended user or the initial transfer to an intermediate person, if used:

(a) Furnish to the intended user and to each person to whom a device is transferred as an intermediary, the following:

(i) A copy of the general license contained in WAC 246-233-020.  If WAC 246-233-020 (3)(b), (c), and (d) or (k) do not apply, those subsections may be omitted;

(ii) A copy of WAC 246-232-050, 246-221-230, 246-221-240, and 246-221-250;

(iii) A list of the services that can only be performed by a specific licensee; and

(iv) Information on acceptable disposal options including estimated costs of disposal;

### REGISTERED GENERAL LICENSEE OBLIGATIONS

- What devices require registration?

WAC 246-233-020(3)(i)Devices containing at least:

| Nuclide | mCi |
| --- | --- |
| Cs-137 | 10 |
| Sr-90 | 0.1 |
| Co-60 | 1 |
| Am-241 | 1 |
| Transuranics: Np, Pu, Cm, Bk, Cf, Es, Fm, Md, No, Lr, etc. | 1 |

- What is the registrant obligated to do?

WAC 246-233-020(3) Any person who owns, acquires, receives, possesses, uses or transfers radioactive material in a device pursuant to the general license in subsection (1) of this section:

(a) Shall assure that all labels affixed to the device at the time of receipt and bearing a statement that removal of the label is prohibited are maintained thereon and shall comply with all instructions and precautions provided by such labels;

(b) Shall assure that the device is tested for leakage of radioactive material and proper operation of the on-off mechanism and indicator, if any, at no longer than six-month intervals or at such other intervals as are specified in the label, however:

(i) Devices containing only krypton need not be tested for leakage of radioactive material; and

(ii) Devices containing only tritium or not more than 100 microcuries of other beta and/or gamma emitting material or 10 microcuries of alpha emitting material need not be tested for any purpose.  Devices held in storage in the original shipping container prior to initial installation need not be tested until immediately prior to use;

(c) Shall assure that the tests required by (b) of this subsection and other testing, installing, servicing, and removing from installation involving the radioactive materials, its shielding or containment, are performed:

(i) In accordance with the instructions provided by the labels; or

(ii) By a person holding a specific license from the department or from the United States Nuclear Regulatory Commission or from any agreement state or from a licensing state to perform such activities;

(d) Shall maintain records showing compliance with the requirements of (b) and (c) of this subsection.  The records shall show the results of tests.  The records also shall show the dates of performance and the names of persons performing, testing, installing, servicing, and removing from installation concerning the radioactive material, its shielding or containment.  Records of tests for leakage of radioactive material required by (b) of this subsection shall be maintained for three years after the next required leak test is performed or the sealed source is transferred or disposed.  Records of tests of the on/off mechanism and indicator required by (b) of this subsection shall be maintained for three years after the next required test of the on/off mechanism and indicator is performed or the sealed source is transferred or disposed.  Records of other testing, installation, servicing, and removal from installation required by (c) of this subsection shall be maintained for a period of three years from the date of the recorded event or until the device is transferred or disposed;

(e) Upon the occurrence of a failure of or damage to, or any indication of a possible failure of or damage to, the shielding of the radioactive material or the on/off mechanism or indicator, or upon the detection of 0.005 microcuries or more removable radioactive material, shall immediately suspend operation of the device until it has been repaired by the manufacturer or other person holding a specific license from the department, the United States Nuclear Regulatory Commission, or from an agreement state or a licensing state to repair such devices, or disposed by transfer to a person authorized by a specific license to receive the radioactive material contained in the device and, within thirty days, furnish to the department a written report containing a brief description of the event and the remedial action taken; and, in the case of detection of 0.005 microcuries or more of removable radioactive material or failure of or damage to a source likely to result in contamination of the premises or the environs, a plan for ensuring that the premises and environs are acceptable for unrestricted use (see WAC 246-246-020);

(f) Shall not abandon the device containing radioactive material;

(g) Except as provided in (h) of this subsection, shall transfer or dispose the device containing radioactive material only by transfer to a person holding a specific license of the department, the United States Nuclear Regulatory Commission, or an agreement state, or a licensing state whose specific license authorizes the person to receive the device and within thirty days after transfer of a device to a specific licensee shall furnish to the department a report containing identification of the device by manufacturer's (or initial transferor's) name, model number, and serial number; the name, address, and license number of the person receiving the device, and the date of transfer.  Prior written approval from the department is required before transferring the device to any other specific licensee not specifically identified in this subsection;

(h) Shall transfer the device to another general licensee only:

(i) Where the device remains in use at a particular location.  In such case, the transferor shall give the transferee a copy of this section, a copy of WAC 246-221-240, 246-221-250, 246-232-050, and 246-232-060, and any safety documents identified in the label of the device and within thirty days of the transfer, report to the department the manufacturer's (or transferor's) name, model number, and serial number of device transferred, the transferee's name and mailing address for the location of use, and the name, title, and phone number of the responsible individual identified by the transferee in accordance with (j) of this subsection to have knowledge of and authority to take actions to ensure compliance with the appropriate regulations and requirements; or

(ii) Where the device is held in storage in the original shipping container at its intended location of use prior to initial use by a general licensee((:));

(i) Shall comply with the provisions of WAC 246-221-240 and 246-221-250 for reporting radiation incidents, theft or loss of licensed material, but shall be exempt from the other requirements of chapters 246-221 and 246-222 WAC;

(j) Shall appoint an individual responsible for having knowledge of the appropriate regulations and requirements and the authority for taking required actions to comply with appropriate regulations and requirements.  The general licensee, through this individual, shall ensure the day-to-day compliance with appropriate regulations and requirements.  This appointment does not relieve the general licensee of any of its responsibility in this regard;

(k)(i) Shall register, in accordance with (k)(i) and (iii) of this subsection, devices containing at least 370 MBq (10 mCi) of Cesium-137, 3.7 MBq (0.1 mCi) of Strontium-90, 37 MBq (1 mCi) of Cobalt-60, or 37 MBq (1 mCi) of Americium-241 or any other transuranic (i.e., element with atomic number greater than uranium (92)), based on the activity indicated on the label.  Each address for a location of use, as described under (k)(iii)(D) of this subsection, represents a separate general licensee and requires a separate registration and fee;

(ii) If in possession of a device meeting the criteria of (k)(i) of this subsection, shall register these devices annually with the department and shall pay the fee required by WAC 246-254-090.  Registration must be done by verifying, correcting, and/or adding to the information provided in a request for registration received from the department.  The registration information must be submitted to the department within thirty days of the date of the request for registration or as otherwise indicated in the request.  In addition, a general licensee holding devices meeting the criteria of (k)(i) of this subsection is subject to the bankruptcy notification requirement in WAC 246-232-050;

(iii) In registering devices, the general licensee shall furnish the following information and any other information specifically requested by the department:

(A) Name and mailing address of the general licensee;

(B) Information about each device:  The manufacturer (or initial transferor), model number, serial number, the radionuclide and activity (as indicated on the label);

(C) Name, title, and telephone number of the responsible person designated as a representative of the general licensee under (j) of this subsection;

(D) Address or location at which the device(s) are used and/or stored.  For portable devices, the address of the primary place of storage;

(E) Certification by the responsible representative of the general licensee that the information concerning the device(s) has been verified through a physical inventory and checking of label information;

(F) Certification by the responsible representative of the general licensee that they are aware of the requirements of the general license;

(iv) Persons generally licensed by the U.S. Nuclear Regulatory Commission, or an agreement state with respect to devices meeting the criteria in (k)(i) of this subsection are not subject to registration requirements if the devices are used in areas subject to Washington state jurisdiction for a period less than one hundred eighty days in any calendar year.  The department will not request registration information from such licensees;

(l) Shall report changes to the mailing address for the location of use (including change in name of general licensee) to the department within thirty days of the effective date of the change.  For a portable device, a report of address change is only required for a change in the device's primary place of storage;

(m) Shall not hold devices that are not in use for longer than two years.  If devices with shutters are not being used, the shutter must be locked in the closed position.  The testing required by (b) of this subsection need not be performed during the period of storage only.  However, when devices are put back into service or transferred to another person, and have not been tested within the required test interval, they must be tested for leakage before use or transfer and the shutter tested before use.  Devices kept in standby for future use are excluded from the two-year time limit if the general licensee performs quarterly physical inventories of these devices while they are in standby.

- What is involved in registering?

WAC246-233-020(3)(k)(iii) In registering devices, the general licensee shall furnish the following information and any other information specifically requested by the department:

(A) Name and mailing address of the general licensee;

(B) Information about each device:  The manufacturer (or initial transferor), model number, serial number, the radionuclide and activity (as indicated on the label);

(C) Name, title, and telephone number of the responsible person designated as a representative of the general licensee under (j) of this subsection;

(D) Address or location at which the device(s) are used and/or stored.  For portable devices, the address of the primary place of storage;

(E) Certification by the responsible representative of the general licensee that the information concerning the device(s) has been verified through a physical inventory and checking of label information;

(F) Certification by the responsible representative of the general licensee that they are aware of the requirements of the general license;

(iv) Persons generally licensed by the U.S. Nuclear Regulatory Commission, or an agreement state with respect to devices meeting the criteria in (k)(i) of this subsection are not subject to registration requirements if the devices are used in areas subject to Washington state jurisdiction for a period less than one hundred eighty days in any calendar year.  The department will not request registration information from such licensees;

(l) Shall report changes to the mailing address for the location of use (including change in name of general licensee) to the department within thirty days of the effective date of the change.  For a portable device, a report of address change is only required for a change in the device's primary place of storage;

- How often do these devices need to be registered?

Annually and pay fees annually.

WAC 246-233-020(3)(k)(ii) If in possession of a device meeting the criteria of (k)(i) of this subsection, shall register these devices annually with the department and shall pay the fee required by WAC 246-254-090.  Registration must be done by verifying, correcting, and/or adding to the information provided in a request for registration received from the department.  The registration information must be submitted to the department within thirty days of the date of the request for registration or as otherwise indicated in the request.

- How do registrants know they need to register these devices, and when do registrants need to register by?

The department will notify registrants that they need to register these generally licensed devices.  The registrant must provide the requested information within 30 days, or as otherwise indicated.

WAC 246-233-020(3)(k)(ii) If in possession of a device meeting the criteria of (k)(i) of this subsection, shall register these devices annually with the department and shall pay the fee required by WAC 246-254-090.  Registration must be done by verifying, correcting, and/or adding to the information provided in a request for registration received from the department.  The registration information must be submitted to the department within thirty days of the date of the request for registration or as otherwise indicated in the request.

- What information needs to be supplied to register a specified generally licensed device?

A registration from, based on the requirements given in WAC 246-233-020(3)(k)(iii), will be provided to the registrant.

- When is the registrant obligated to contact the Department?

When reporting radiation incidents, bankruptcy, failure of the device on/off mechanism or shielding, a positive leak test result, or the transfer or disposal of the device.  The registrant in NOT required to contact the department when receiving a new device.

- What other regulations is the registrant subjected to?

| WAC |  |
| --- | --- |
| 246-220-020 | General Provisions: Records |
| 246-220-030 | Inspections. |
| 246-220-040 | Tests and surveys. |
| 246-220-060 | Violations. |
| 246-220-070 | Impounding. |
| 246-220-100 | Additional requirements. |
| 246-221-240 | Radiation Protection Standards: Reports of stolen, lost or missing radiation sources. |
| 246-221-250 | Notification of incidents |
| 246-232-050 | Radioactive Material -- Licensing Applicability: Terms and conditions of licenses. |
| 246-232-060 | Termination of licenses and decommissioning of sites and separate buildings or outdoor areas. |
| 246-232-070 | Modification and revocation of licenses. |
| 236-232-080 | Transfer of material. |
| 246-232-090 | Transportation. |

### REGISTRATION PROCESS

- How does a registrant receive a registration form?

The department will send the registrant a Registration Packet, that includes the Registration Form.  Registrants that are required to register their devices will be determined from the vendor quarterly reports.

- What will the registrant receive at the end of the registration process?

A State of Washington Radioactive Material General License Registration that will detail conditions on how the device shall be handled, and servicing, record keeping and reporting requirements.

- Will the State of Washington Radioactive Material General License Registration have an expiration date?

Yes, this date is used to track the annual registration requirement.

WAC 246-233-020(3)(k)(ii) If in possession of a device meeting the criteria of (k)(i) of this subsection, shall register these devices annually with the department and shall pay the fee required by WAC 246-254-090.  Registration must be done by verifying, correcting, and/or adding to the information provided in a request for registration received from the department.  The registration information must be submitted to the department within thirty days of the date of the request for registration or as otherwise indicated in the request.

### DISPOSAL AND TRANSFERS

- How may the registrant dispose of the registered device?

WAC 246-233-020(3)(g) Except as provided in (h) of this subsection, shall transfer or dispose the device containing radioactive material only by transfer to a person holding a specific license of the department, the United States Nuclear Regulatory Commission, or an agreement state, or a licensing state whose specific license authorizes the person to receive the device and within thirty days after transfer of a device to a specific licensee shall furnish to the department a report containing identification of the device by manufacturer's (or initial transferor's) name, model number, and serial number; the name, address, and license number of the person receiving the device, and the date of transfer.  Prior written approval from the department is required before transferring the device to any other specific licensee not specifically identified in this subsection;

- May the registrant transfer the registered device to another general licensee?

WAC 246-233-020(3)(h) Shall transfer the device to another general licensee only:

(i) Where the device remains in use at a particular location.  In such case, the transferor shall give the transferee a copy of this section, a copy of WAC 246-221-240, 246-221-250, 246-232-050, and 246-232-060, and any safety documents identified in the label of the device and within thirty days of the transfer, report to the department the manufacturer's (or transferor's) name, model number, and serial number of device transferred, the transferee's name and mailing address for the location of use, and the name, title, and phone number of the responsible individual identified by the transferee in accordance with (j) of this subsection to have knowledge of and authority to take actions to ensure compliance with the appropriate regulations and requirements; or

(ii) Where the device is held in storage in the original shipping container at its intended location of use prior to initial use by a general licensee;

(i) Shall comply with the provisions of WAC 246-221-240 and 246-221-250 for reporting radiation incidents, theft or loss of licensed material, but shall be exempt from the other requirements of chapters 246-221 and 246-222 WAC;